Experimental tests results

The results obtained by the experimental tests of aerosol leakage (P) through the filtering material of the filtering half masks and the resistance which they provide at various aerosol flow rates (p), depending on the exposure time duration, are all shown in Tables 1 and 2. Presented results are the average values of nine parallel measurements, with variation less than 30%.

TABLE 1. The leakage values of the filtering material samples depending on the time, at the test aerosol flow of 30 and 95 dm3 / min (indicated by an asterisk)

| Time (min) | P (%) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | EM1 | EM1* | EM2 | EM2* | HM1 | HM1* | HM2 | HM2* | HM3 | HM3* |
| 0 | 1.1773 | 0.8557 | 1.2192 | 0.9198 | 0.1783 | 0.0686 | 0.1331 | 0.0612 | 1.3636 | 1.2347 |
| 5 | 1.1511 | 0.8469 | 1.1905 | 0.9090 | 0.1691 | 0.0630 | 0.1278 | 0.0563 | 1.3224 | 1.1811 |
| 10 | 1.1163 | 0.8278 | 1.1245 | 0.8852 | 0.1533 | 0.0552 | 0.1189 | 0.0487 | 1.2678 | 1.1587 |
| 15 | 1.0837 | 0.8093 | 1.0985 | 0.8697 | 0.1376 | 0.0463 | 0.1027 | 0.0429 | 1.2014 | 1.1142 |
| 20 | 1.0722 | 0.7205 | 1.0753 | 0.7345 | 0.1271 | 0.0434 | 0.0958 | 0.0395 | 1.1811 | 1.0965 |
| 25 | 1.0598 | 0.7048 | 1.0526 | 0.6450 | 0.1041 | 0.0400 | 0.0836 | 0.0382 | 1.1587 | 1.0777 |
| 30 | 1.0378 | 0.6883 | 1.0202 | 0.7094 | 0.0897 | 0.0385 | 0.0703 | 0.0371 | 1.1142 | 1.0398 |

TABLE 2. The resistance values of filtering material samples depending on the time, at the test aerosol flow of 30 and 95 dm3 / min (indicated by an asterisk)

| Time (min) | p (Pa) |  |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
|  | EM1 | EM1* | EM2 | EM2* | HM1 | HM1* | HM2 | HM2* | HM3 | HM3* |
| 0 | 50 | 140 | 55 | 130 | 75 | 260 | 80 | 270 | 15 | 140 |
| 5 | 60 | 140 | 55 | 130 | 75 | 300 | 80 | 310 | 15 | 150 |
| 10 | 75 | 150 | 60 | 140 | 85 | 340 | 90 | 360 | 20 | 160 |
| 15 | 85 | 150 | 60 | 140 | 120 | 370 | 130 | 380 | 20 | 180 |
| 20 | 90 | 160 | 65 | 150 | 140 | 390 | 140 | 400 | 35 | 200 |
| 25 | 100 | 170 | 70 | 150 | 160 | 400 | 170 | 410 | 40 | 250 |
| 30 | 110 | 180 | 80 | 170 | 190 | 410 | 210 | 420 | 60 | 330 |